Impact of Naspharyngeal Swab Types on Bordetella pertussis Detection by PCR and Culture.

J. Cloud1, W. Hymas2, and K. C. Carroll1,3

1ARUP Institute for Clinical and Experimental Pathology, Salt Lake City, UT.

2ARUP Molecular Infectious Diseases Laboratory, Salt Lake City, UT.

3University of Utah, Department of Pathology, Salt Lake City, UT.

ABSTRACT

B. pertussis may be diagnosed by direct fluorescent antibody staining (DFA), culture or PCR of nasopharyngeal specimens.   Specimens submitted for culture may be collected using a variety of different swabs made of calcium-alginate, dacron or rayon. Although alginate is known to inhibit PCR, calcium-alginate swabs have been recommended for culture of nasopharyngeal samples.   Dacron and rayon tipped swabs do not inhibit PCR, but the ability to which they preserve B. pertussis organisms for culture remains a question. A study was designed to assess the performance of various swabs and transport media routinely used to collect specimens submitted for Bordetella culture and/or PCR.  The swab products evaluated included: calcium-alginate, dacron, and rayon nasopharyngeal swabs, each with and without transport media (Regan-Lowe or Amies with charcoal).  Serial 10-fold dilutions from a 0.5 McFarland suspension of organisms in saline were used to inoculate swabs.  Culture was performed on Regan-Lowe media using standard laboratory protocols.  A LightCycler PCR assay with an internal inhibition control was used in the PCR assay for all samples.  Of 4 separate sets of dilutions, all of the calcium-alginate swabs were inhibited in the PCR reaction without using a swab pretreatment protocol.  Three swab pretreatment protocols were evaluated with 6 additional sets of dilutions resulting in 100%, 71% and 21% inhibition of calcium alginate swabs.  No inhibition was detected in any PCR reactions with DNA from dacron or rayon swabs, with or without transport media.  Dacron, rayon, and calcium alginate swabs performed similarly for recovery of B. pertussis in culture.  The Amies with charcoal transport system performed poorly for culture.  Calcium alginate swabs should not be used for PCR-based detection of Bordetella pertussis whereas dacron or rayon swabs are an excellent choice for both PCR and culture.

INTRODUCTION

Recent articles describing assays of real-time PCR for the detection of Bordetella species have not resolved the problem of calcium-alginate inhibition (Sloan et al, 2002).  The inhibitory effects of calcium-alginate tipped nasopharyngeal (NP) swabs have been reported by Wadowsky et al. (1994).  Even with this knowledge, laboratories offering PCR for the detection of B. pertussis continue to accept samples collected with calcium-alginate swabs.  Admittedly, it is convenient to use a single swab for both culture and PCR.  However, evaluations of pertussis PCR assays compared with culture (Loeffelholz et al, 1999, Sloan et al., 2002) may be misleading if the samples were collected with calcium-alginate swabs.

It is not practical to collect two swabs per patient; i.e., one calcium-alginate for culture and one synthetic for PCR, as suggested by Wadowsky et al. (1994).  Some laboratories may be concerned about the use of synthetic fiber swabs, such as dacron or rayon, for culture.  Nasopharyngeal specimens are routinely collected on calcium-alginate swabs and submitted to the lab for detection of B. pertussis by culture.  Early reports suggested calcium-alginate swabs were superior to dacron, rayon, or cotton wool in their ability to maintain viability of B. pertussis (Hoppe and Weiss, 1987).  Additional reports have supported the fact that cotton wool swabs perform poorly for isolation of B. pertussis (Hoppe and Weiss, 1987, Hoppe et al., 1986, Ross and Cumming, 1981).  It has been suggested that fatty acids on some cotton wool swabs inhibit bacterial growth (Hoppe and Weiss, 1987).

We compared several types of swabs to determine which single material can optimally be used for both culture and PCR.  Calcium-alginate, dacron, and rayon tipped nasopharyngeal swabs, with or without Regan-Lowe transport media, were evaluated for their ability to maintain viable B. pertussis as well as their effect on detection by PCR.  Rayon swabs in the presence of Amies with charcoal transport medium (COPAN, Corona, Calif.) as well as the DNA/RNA Protect system (Sierra Molecular Sonora, Ca.) were included in the study.   All swabs were produced by the same manufacturer in effort to standardize the size of the swab head.

MATERIALS AND METHODS

Bacterial Strains.  Ten clinical isolates of B. pertussis previously identified by Primary Childrens Medical Center (Salt Lake City, Utah) were utilized in the swab/transport evaluation.

Sample Preparation.  Bacterial suspensions were prepared to match a 0.5 McFarland standard.  Dilutions from which 1:103, 1:104, 1:105, 1:106, and 1:107 were made for these analyses.  Five microliters of each dilution were inoculated onto each type of swab in triplicate. One swab of each triplicate was held at room temperature for 24 hours and analyzed by PCR only.  The remaining two swabs were stabbed into Regan-Lowe Deep media (Becton Dickinson Microbiology Systems, Sparks, MD) and incubated at 36oC for 48 hours.  One of the inoculated swabs incubated in the agar tubes was analyzed by PCR while the other was set up for culture.

Culture.  Growth controls were prepared to determine if cultivable organisms were present in the dilutions.  Ten samples of each dilution were assessed for each type of swab.  Swabs to be cultured were rolled onto a quadrant of a Regan-Lowe plate, then streaked out into the remaining quadrants.  The plates were incubated at 36oC and protected from desiccation.  Plates were examined daily for 7 days.  Any colonies detected at 7 days were deemed positive.

Preparation of swabs for nucleic acid extraction.  Swabs were prepared for DNA extraction by one of 3 protocols.  The first protocol involved direct DNA extraction in which swabs were immersed in 200 l of tissue lysis buffer (ATL) supplied in the QIAamp DNA Mini Kit (Qiagen, Valencia, Calif.).  Swabs were vortexed, wrung-out and discarded.  The remaining solution was used for nucleic acid extraction.  In the second protocol swabs were immersed in 1 ml of phosphate buffered saline (Invitrogen, Carlsbad, Calif.), incubated for 10 minutes at room temperature, vortexed, wrung-out and discarded.  The remaining suspension was centrifuged for 10 minutes at 6,000 x g, the supernatant decanted, and the pellet used for nucleic acid extraction.  The third protocol differed from the second by soaking swabs in 2 ml of phosphate buffered saline (PBS) for 10 minutes and, after vortexing, using 1 ml of this solution for centrifugation and retention of the pellet for extraction.

Nucleic acid extraction for PCR.  Nucleic acid was extracted from swab preparations or 5 l of bacterial dilutions with the QIAamp DNA Mini Kit according to manufacturer’s instructions and eluted from the column in 100 l of TE prime buffer (10 mM Tris, 0.1 mM EDTA, pH 8.0).  Extracted DNA was directly analyzed by LightCycler PCR without storage.

Primers and probes.  Primers and probes for Bordetella real-time PCR have been described previously (Reischl et al., 2001).  Oligonucleotides were synthesized by IT BioChem (Salt Lake City, Utah).  All primers and probes were resuspended using TE prime buffer.

PCR Inhibition Control.  PCR inhibition was assessed by the addition of approximately 50 fg of previously extracted B. pertussis DNA.  The DNA was applied in 5 l volumes to each well of microtiter plates, allowed to desiccate, and stored at –20oC until put into use.  Once thawed for use, the inhibition control plates were held at room temperature for up to 1 month.  An aliquot of each PCR reaction mixture was applied to a well of the inhibition control plate before analysis (described below).

LightCycler PCR.  A multiplex real-time PCR assay was used for PCR analysis of Bordetella sp. (presented at the 102nd General Meeting of the American Society for Microbiology, Salt Lake City, Utah, 19 to 23 May 2002).  The PCR master mix consisted of 3 mM MgCl2, 1X LightCycler FastStart DNA Master Hybridization Probes (Roche Diagnostics, Indianapolis, Ind.) prepared according to manufacturer’s instructions, 0.5 M each forward primer, 1.0 M each reverse primer, 0.2 M B. pertussis anchor and reporter probes, 1.0 M B. parapertussis anchor and reporter probes, and 20 U/ml uracil-N-glycosylase (UNG, Applied Biosystems, Foster City, Calif.).  The master mix was prepared to achieve a total PCR volume of 30 l per sample (24 l master mix plus 6 l sample DNA or control).  Twenty microliters were transferred to a LightCycler capillary (Roche Diagnostics, Indianapolis, Ind).  The remaining 10 l were mixed in a well of the inhibition control plate (described above) prior to being transferred to a capillary tube.  All capillary tubes were capped, inserted into a LightCycler carousel and centrifuged prior to amplification and detection. The cycling parameters began with 1 cycle of 40oC for 4 minutes (activation of UNG) and 95oC for 9 minutes (inactivation of UNG and activation of FastStart polymerase).  The second part of the cycling protocol was 50 cycles of 95oC for 1 s, 53oC for 8 s, and 72oC for 18 s.  The last portion of the protocol was a cool-down step at 40oC for 5 s.  Data analysis was performed with version 3.5.3 LightCycler software.

RESULTS

PCR.  Dilutions of 10 separate B. pertussis isolates in suspension were analyzed.  Using dry swabs inoculated with a 10-3 dilution of organisms from a 0.5 McFarland suspension, PCR was positive for all inoculated dacron and rayon swabs.  However, for calcium alginate, PCR was only positive for 50% of all samples analyzed at the same dilution (table 1).  At the 10-4 dilution, PCR positivity was 70% and 80% for dacron and rayon swabs, respectively, while only 20% were positive with calcium-alginate swabs.  As a control, the same volume of each dilution was analyzed and resulted in 100% positivity at the 10-3 and 10-4 dilutions, 80% at the 10-5 dilution, and no positivity at the 10-6 dilution.

In an effort to mimic the recommended transport system for swab samples submitted for culture (Hoppe et al., 1986), inoculated swabs were inserted into Regan-Lowe deep transport media and incubated at 36oC for 48 hours before testing.  Table 2 describes the PCR results of swabs inserted into transport media.  Dacron and rayon swabs both resulted in 100% positivity with samples at both the 10-4 and 10-5 dilutions, while calcium-alginate swabs were positive only 20% of the time at the 10-5 dilution.  The 10-4 dilution was positive with calcium-alginate in only 40% of the samples.  Because of our anticipation for bacterial growth during the 48-hour incubation, the 10-3 dilution was not tested with swabs inserted into Regan-Lowe media.  Rayon swabs inserted into Amies with charcoal medium, rather than Regan-Lowe, only showed 80% positivity in the 10-4 dilution.

A newer swab transport system for PCR is the DNA/RNA Protect system consisting of a standard sized rayon swab with a plastic shaft which is inserted into a tube containing a proprietary moist sponge-like material.  PCR positivity was 100% at the 10-3 dilution, while dropping to 90% and 50% at the 10-4 and 10-5 dilutions, respectively.

Culture.  Culture of dacron and rayon swabs after incubation in Regan-Lowe transport media for 48 hours prior to plating was just as good or better than with calcium-alginate swabs (table 3).  Dacron swabs yielded 40% recovery at the 10-6 dilution compared to 20% recovery from rayon and calcium-alginate swabs. All three types of swabs showed 100% recovery at the 10-4 dilution.  Both rayon and dacron showed 90% recovery while calcium-alginate showed a 60% recovery at the 10-5 dilution.  Rayon swabs inserted into Amies with charcoal media supported growth of B. pertussis only 20% of the time at the 10-3 dilution.  No growth was detected at the 10-4 or 10-5 dilutions using the Amies with charcoal transport media.

PCR Inhibition.  PCR inhibition was only detected when calcium-alginate swabs were used.  Inhibition rates were determined using 3 separate protocols for sample preparation prior to DNA extraction (table 4).  The lysis buffer preparation resulted in 100% inhibition using calcium-alginate swabs.  When soaking the swabs in 1 ml of PBS prior to pelleting and extraction the inhibition rate dropped to 71%.  Better performance was seen by soaking the swabs in 2 ml of PBS and pelleting half of the 2 ml suspension prior to extraction which dropped the inhibition rate to 21%.

DISCUSSION

Dacron or rayon NP swabs with aluminum wire shafts can and should be used for PCR for the detection of B. pertussis.  Also, if the swab is to be plated for culture prior to PCR, dacron or rayon swabs can be used as well.  Neither type of synthetic swab, dacron and rayon, is superior to the other according to our studies.  They both perform well for PCR and culture.    For laboratories aiming to validate a PCR test for B. pertussis detection, sample retrieval would be easy if the same swab submitted in Regan-Lowe media could be used.  The swab samples would have been incubated for 48 hours prior to plating, after which remaining organisms would be available on the swab for PCR.

As PCR becomes more popular for detection of infectious diseases, attention must be given to collection systems.  The DNA/RNA Protect system consists of a swab with a moist sponge to preserve the nucleic acid without diluting out the sample, showed promising results in our study.  The nucleic acid preservative allows the bacteria to be centrifuged into a pellet.  Although the swabs included in the DNA/RNA Protect system for our study were rayon swabs with standard sized heads and plastic shafts, NP swabs with aluminum wire shafts can be used with this system as well.

The main component of calcium-alginate swabs responsible for PCR inhibition has been reported to be the alginate, which is a crude extract from seaweed (Wadowsky et al., 1994).  Reports of whether or not components of the aluminum shaft are inhibitory have been conflicting.  Wadowsky et al. (1994) reported inhibition due to the aluminum shaft where Rasmussen et al. (1992) and He et al. (1993) had success using aluminum shafts in their PCR assays.  We did not see any inhibition with swabs containing aluminum shafts, even after we have soaked the swabs, with 2 inches of aluminum shaft, overnight before extraction and analysis of DNA (personal observation).  The controversy may be due to differences in swab components provided by manufacturers or to differences in extraction methods.

Unlike Regan-Lowe medium, Amies with charcoal supports growth of B. pertussis poorly and therefore should not be used to increase sensitivity of detection. This finding is in agreement with other reports (Ross and Cumming, 1981, Hoppe et al., 1986).  Media containing animal blood (i.e. Regan-Lowe media) is known to be advantageous to the support of growth of B. pertussis (Hoppe and Schlagenhauf, 1989). Amies medium does not contain blood.  If a sample for culture of B. pertussis is collected with the Amies with charcoal transport system, immediate plating is more likely to be better than incubation for 48 hours, thus improving the turn around time.  For PCR detection, the Amies with charcoal transport system performed similar to PCR of dry swabs as noted in tables 1 and 2.  In an effort to curb costs, dry swabs may well be used for PCR.

In conclusion, if seeking a single swab that can be used for either B. pertussis culture or PCR, rayon or dacron swabs are recommended.  Calcium-alginate swabs should not be used for the PCR detection of B. pertussis. If a calcium-alginate or an unknown type of swab is received for PCR detection of B. pertussis, washing the swab in 2 ml of PBS reduces the inhibition rate due to calcium-alginate.   It is strongly advised to use only dacron or rayon swabs for all NP swab collections for culture and/or PCR of B. pertussis.

REFERENCES

- He, Q., J. Mertsola, H. Soini, M. Skurnik, O. Ruuskanen, and M. K. Viljanen.  1993.  Comparison of polymerase chain reaction with culture and enzyme immunoassay for diagnosis of pertussis.  J. Clin. Microbiol.  31:642-645.

- Hoppe, J. E. and A. Weiss.  1987.  Recovery of Bordetella pertussis from four kinds of swabs.  Eur. J. Clin. Microbiol.  6:203-205.

- Hoppe, J. E., S. Worz, and K. Botzenhart.  1986.  Comparison of specimen transport systems for Bordetella pertussis.   Eur. J. Clin. Microbiol.  5:671-673.

- Hoppe, J. E. and M. Schlagenhauf.  1989.  Comparison of three kinds of blood and two incubation atmospheres for cultivation of Bordetella pertussis on charcoal agar.  J. Clin. Microbiol.  27:21152117.

- Loeffelholz M. J., C. J. Thompson, K. S. Long, M. J. R. Gilchrist.  1999.  Comparison of PCR, culture, and direct fluorescent antibody testing for detection of Bordetella pertussis.  J. Clin. Microbiol.  37:2872-2876.

- Rasmussen S. J., F. P. Douglas, and P. Timms.  1992.  PCR detection and differentiation of Chlamydia pneumonia, Chlamydia psittaci and Chlamydia trachomatis.  Mol. Cell. Probes.  6:389-94.

- Reischl, U., K. Kösters, B. Leppmeier, H.-J. Linde, and N. Lehn.  2001.  Rapid detection and simultaneous differentiation of Bordetella pertussis and Bordetella parapertussis in clinical specimens by LightCycler PCR.  P. 31-44. In: U. Reischl, C. Wittwer, and F. Cockerill (Eds), Rapid cycle Real-Time PCR: Methods and Applications, Springer Press, Heidelberg.

- Ross, P. W. and C. G. Cumming.  1981.  Isolation of Bordetella pertussis from swabs.  Br. Med. J.  283:403-404

- Sloan, L. M., M. K. Hopkins, P. S. Mitchell, E. A. Vetter, J. E. Rosenblatt, W. S. Harmsen, F. R. Cockerill, and R. Patel.  2002.  Multiplex LightCycler PCR assay for detection and differentiation of Bordetella pertussis and Bordetella parapertussis in nasopharyngeal specimens.  J. Clin. Microbiol.  40:96-100.

- Wadowsky R. M., S. Laus, T. Libert, S. J. States, and G. D. Ehrlich.  1994.   Inhibition of PCR-based assay for Bordetella pertussis by using calcium alginate fiber and aluminum shaft components of a nasopharyngeal swab.  J Clin Microbiol.  32:1054-1057.

Table 1.  Percent of B. pertussis detected by PCR using various types of swabs.  As a reference, dilutions were also analyzed without a swab.

| Dilution | No swab (control) | Dacron | Rayon | Calcium-alginate | DNA/RNA Protect |
| --- | --- | --- | --- | --- | --- |
| 10-3 (n=10)* 10-4 (n=10) 10-5 (n=10) 10-6 (n=10) | 100% 100% 80% 0% | 100% 70% 10% 0% | 100% 80% 10% 0% | 50% 20% 10% 0% | 100% 90% 50% 0% |

* Value in parenthesis represents number of isolates tested at that dilution.

Table 2.  Percent of B. pertussis detected by PCR using various types of swabs incubated for 48 hours in Regan-Lowe or Amies with charcoal media.

| Dilution | Dacron/Regan-Lowe | Rayon/Regan-Lowe | Calcium-alginate/Regan-Lowe | Rayon/Amies with Charcoal |
| --- | --- | --- | --- | --- |
| 10-3 (n=10)* 10-4 (n=10) 10-5 (n=10) 10-6 (n=10) 10-7 (n=10) | N/P 100% 100% 50% 0% | N/P 100% 100% 20% 0% | N/P 40% 20% 0% 0% | 100% 80% 40% 0% N/P |

N/P = not performed.

* Value in parenthesis represents number of isolates tested at that dilution.

Table 3.  Percent of swabs yielding B. pertussis by culture.

| Dilution | Dacron/Regan-Lowe | Rayon/Regan-Lowe | Calcium-alginate/Regan-Lowe | Rayon/Amies with Charcoal |
| --- | --- | --- | --- | --- |
| 10-3 (n=10)* 10-4 (n=10) 10-5 (n=10) 10-6 (n=10) 10-7 (n=10) | N/P 100% 90% 40% 0% | N/P 100% 90% 20% 0% | N/P 100% 60% 20% 0% | 20% 0% 0% N/P N/P |

N/P = not performed.

* Value in parenthesis represents number of isolates tested at that dilution.

Table 4.  Percent inhibition of B. pertussis PCR of DNA samples from various types of swabs after pretreatment by one of three protocols.

| Swab Pretreatment* | Dacron | Rayon | Calcium-alginate | DNA/RNA Protect (rayon) | Amies with Charcoal (rayon) |
| --- | --- | --- | --- | --- | --- |
| Direct Lysis Buffer Rinse in 1 ml PBS Rinse in 2 ml PBS | 0% (n=32) 0% (n=24) 0% (n=24) | 0% (n=32) 0% (n=24) 0% (n=24) | 100% (n=32) 71% (n=24) 21% (n=24) | 0% (n=16) 0% (n=12) 0% (n=12) | 0% (n=16) 0% (n=12) 0% (n=12) |

* See text for explanation.